## Slide 1: 2003 eVent advanced ventilator
The 2003 eVent Inspiration was a low cost concept that incorporates all modes of pressure and volume control ventilation. Its pneumatic diagram is on page 17 of its published service manual:
http://www.frankshospitalworkshop.com/equipment/documents/ventilators/service_manuals/eVent_Medical_Inspiration_Ventilator_-_Service_manual.pdf
Voice coil actuator
0.02L/min flow to keep water droplets and patient cough secretion out of the pressure measurement lines.
Compressor for
transport or supply fail ventilation
0.02L/min flow past oxygen cell.
FR
FR
FR
Safety relief
Reproduced from a published document; but without express permission.

## Slide 2: Equivalent blender design
Ptr3
NR3 22mm
PR2
SV2
NC
SV3
NC
SV4
NC
FR1
FR2
Gas Output
Flow rate 60 L/min
4mbar to 35mbar
22mm conical connector, Male
Oxygen
2.5 to 5bar
NR4 22mm
Ptr4
Pressure monitor
Gas Return
Exhaust
Exhaust
Ptr2
PR1
SV1
NC
Air
2.5 to 5bar
Safety shutdown
(monitor)
Pop-off safety relief
(1.5bar)
500ml
Mixing
Chamber
(max 1 bar)
Blender control
Ptr1
By further making SV3 and SV4 proportional, and adding a proximal flow sensor, the design is effectively similar to an eVent Inspiration and can perform all advanced Volume Controlled modes of ventilation
Oxygen cell
Mixing chamber must not exceed the pressure-volume threshold for becoming classed as a pressure vessel. Such vessels can store a lot of energy and therefore have strict, complicated safety regulations.
Remember, the breathing circuit inspiratory tube tends to have a volume of about 1.5L, in which gas flows faster in the centre of the tube and slower on the wall. Mixing will therefore continue throughout the 1.5L tube towards the patient.
Differs from the eVent Inspiration in being continuous and not pressure cycled.
NR1
NR2
NR1, 2 guards against reverse flow into the hospital supply lines.
EV1

## Slide 3: Blending cycle
Flow sensor counts through a pre-determined number of air units, followed by a number of oxygen units. The number of units corresponds to a volume. Because PR1 and PR2 are set to the same pressure, there is no need to linearize and convert the volume units into a SI unit.
The blend is determined by the proportional switch point between the two gas units. The pre-determined unit count thresholds are set, in software, so that the blender cycles approximately once every breath (about every 4 seconds).
Remember, oxygen is 3% denser than air, and FiO2 is a volume/volume measure. PTR1 is a mass-flow sensor. The oxygen unit might need to be increased by 3% for volumetric proportioning. Test and see if a small correction factor is needed. The international standard set a tolerance of 3%.
The pressure in the mixing chamber equals the input regulators PR1 and PR2. It must be 1bar or less – or else pressure vessel controls become necessary in many parts of the world. This calls for SV3 having a larger orifice, to secure the flow rate.
Say, PTR1 is sampled at 10-bit (0 to 1024) every 2mS, and the PTR1 flow orifice creates an ADC integer value of 800 when flow is 60L/min. Each ADC value of 800 represents a volume of 2ml.
Over a full 500ml, 3s breath cycle, the average ADC value is 133 (because there is no flow from ventilator during the inspiration plateau and exhalation phase). 1500 samples over 3s add up to a long integer value of 200,000. This should be the blender cycle reset point.
The variable switch value between 0 and 200,000 creates an oxygen concentration. In the above graph, the switch point is set to 100,000, producing a mix of 50% by volume of oxygen and 50% by volume of air. The resulting FiO2 = (50% x 100% O2) + (50% x 21% O2) = 60.5%.
Reset cycle
Switch gas
Sum of ADC values
time
0
Air intake
Oxygen intake

## Slide 4: Pseudo code for blender control
// Declared constants
MAX_VALUE = 200000
- - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - -
// User interface
IF
FiO2 value is changed // An integer value between 21 and 100
SWITCH_VALUE = (-2533 * FiO2) + 253300
ENDIF
- - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - -
// Routine called every 2mS
Read ADC for FLOW_VALUE // PTR1
VOLUME = VOLUME + FLOW_VALUE
IF VOLUME > MAX_VALUE
Close SV2 // Oxygen valve
Open SV1 // Air valve
VOLUME is zero // Reset cycle
ELSE
IF VOLUME >= SWITCH_VALUE
Open SV2 // Oxygen valve
Close SV1 // Air valve
ENDIF
ENDIF
MAX_VALUE
SWITCH_VALUE
VOLUME
time
0
Air intake
Oxygen intake