## Slide 1: 3 suggested initial sprints
Controller
Monitor
Mechanical layout

## Slide 2: Controller
Controller electronics
Ptr1
2m compressed gas line (8mm PU hose connected directly to pressure regulator)
PR
SV2
NC
SV3
NC
Receives pre-blended gas
Breadboard
External PSU
13.5v
Mains power
EV1
U2
FR1
FR2
CN4
CN5
CN7
CN8
CN11
CN12
CN13
CN15
TB1
TB2
CN9
CN10
5L plastic bottle, or (better) a test lung
Prove that the pressure wave can be reproduced
BPM 10 to 30
EP 4mbar to 20mbar
IP 4mbar to 35mbar
IP
Ti
Te
EP
Pressure
time
Ti = 60/BPM/3 seconds
Te = Ti x 2 seconds

## Slide 3: Controller – pressure wave shape
At BPM = 10: Ti = 2s
At BPM = 30: Ti = 0.7s
Tde should be as short as possible. The actual Tde depends on the lungs natural elasticity and resistance in the breathing circuit. As long as it looks reasonable sharp then the CO2 washout will happen.
IP
Tdi
Tde
EP
Pressure
time
IP
EP
Pressure
time
3 waveform examples:
They all deliver the same amount of ‘work’ and fill the lung with an equivalent volume of gas.
#1 uses the least pressure (which is good), but the abrupt rise is faster than the lung can comply. This risk putting shear stress on the lung tissue.
#3 produces a gentler rise (which is good), but it requires a higher, shorter plateau (which is bad). Lung tissue is sensitive to pressure (barotrauma).
#2 is the best compromise. The plateau should be about 2/3 of the Ti time – meaning Tdi should be no more than 1/3 of the rise time. With the fastest Ti being 0.67s, the best fixed Tdi = 0.2s.
A plastic bottle test lung tends to be less compliant and fills more straightforwardly (compared to the intricate lung). The testing should focus on Tdi = 0.15s.
Tdi is adjusted by regulator PR, or by changing the orifice size in SV2.
3
2
1
IP
EP
Pressure
time
Small over-/under-shoot spikes are tolerated, but undesirable
Large and long over/-under-shots are not tolerated

## Slide 4
SV3
1v
~0.1v
10k
1k
1 bar
700 mbar
Valve
orifice
Ambient
(+0 mbar)
~50 mbar
Flow resistive
orifice 1
Flow resistive
orifice 2
Diaphragm must resist up to 35 mbar patient pressure, but acting on a smaller surface.
Pressure here acts on a bigger surface and it could be less than 35 mbar, but use about 50 mbar for safety margin. Do not exceed 100mbar.
There are similarity to an electric circuit, but they are not entirely the same.
It is ok to conceptualise in the same way.

## Slide 5: Monitor
Monitor electronics
Ptr2
16 segment LCD
Mute
History
SV1
NC
Breadboard
External PSU
13.5v
Mains power
Audible alarm
Visible alarm
U1
SPK
First:
Use a set of test data in a memory array for developing the:
Breath detection (later copied to the Controller)
Tdi and Tde calculations
Over-pressure >40mbar detection and shutdown (later copied to and repeated in Controller for duality).
Second:
Verify with an electrical simulation signal (e.g DAC from a MPU), if time permits. This provides a controlled/noise-free test signal.
Third:
Test on real breathing person (use an AMBU bag if available). This provides a normal noise environment.
IP
Tdi
Tde
EP
time
First
Second
Third
Patient effort creates small drop
Patient circuit pressure
time
tf = 50mS
ts = 2000mS
offset
Breath detect
Breath onset

## Slide 6: Mechanical
Identifying part and organising layout.
Design for Manufacturing.
Use rubber grommets and rubber feet to dampen the mechanical transfer of the solenoid valves’ ‘click’ noise (which are irritating and a source of anxiety in patients).