## Slide 1: FiO2 sensor addition
Ptr1
NR1 22mm
PR
SV1
NC
SV2
NC
SV3
NC
FR1
FR2
Oxygen 2.5 - 5bar
NR2 22mm
Ptr2
FR3
O2 cell
Blender
Air 2.5 - 5bar
PR is not required if the Blender incorporates a pressure regulator
https://www.maxtec.com/product/sensing/fio2/max-250/
Cell reacts with oxygen, producing a voltage of about 13.5mV in ambient air
FR3 set to produce about 20mL/min sampling flow
Oxygen cell signal conditioning, to bring it into ADC range
The cell has a linear response (deviates less than 1%) over the full 21% to 100% range.
The cell depletes over 2 – 4 years (depending on amount of oxygen exposure). The cell voltage drifts down during this time.
The cell needs calibrating at least monthly, when in use. Calibration steps:
Flow pure air, allow 30 sec to settle, measure the 21% value. If below 10mV, then warn about need to change cell (maintain low priority alarm, but allow ventilator use until 9mV).
Flow pure oxygen, allow to settle for at least 30 sec, measure the 100% value.
The 2 values are used to calculate the FiO2 from the cell voltage.